## Document_Control
Document Control
Document Title | Design Controls — SpO2 AI Eval Pipeline
Document Version | 1.0
Device/Software Version | SpO2 AI Eval Pipeline v2.1
Date Generated | 2026-04-28
Safety Class | C
QMS Document Number
SOP Reference
Related Risk Management File
Approval Signatures
Role | Name | Date | Signature
Prepared By
Reviewed By (QA)
Approved By (Management)

## User_Needs
ID | Need Statement | Source | Priority | Acceptance Criteria
UN-001 | System shall classify SpO2 traces into urgency tiers (normal, borderline, urgent, emergency) | Clinical Requirements | High | Classification accuracy >= 85% across all tiers on held-out test set
UN-002 | System shall detect desaturation events below GA-adjusted thresholds | Neonatal Guidelines (Castillo 2008) | High | Sensitivity >= 90% for desaturation events; false negative rate < 5%
UN-003 | System shall generate structured handoff reports for clinical review | Nurse Workflow Analysis | Medium | Reports contain all required fields per SBAR format; 95% nurse satisfaction
UN-004 | System shall flag motion artifacts using accelerometer correlation | Signal Processing Requirements | Medium | Artifact detection precision >= 80%; recall >= 75%
UN-005 | System shall route ambiguous cases to expert queue with confidence scores | Clinical Safety Assessment | High | All cases with confidence < 0.7 routed to expert queue; zero missed critical cases

## Design_Inputs
ID | Linked UN | Requirement | Type | SW Safety Class | Safety Class Rationale
DI-001 | UN-001 | Implement 4-tier classification (normal >= 95%, borderline 90-94%, urgent 85-89%, emergency < 85%) with GA-adjusted thresholds | Functional | C | Class C: Controls desaturation detection; misclassification could delay clinical intervention causing serious harm
DI-002 | UN-002 | Detect SpO2 drops >= 4% from baseline sustained for >= 10 seconds, adjusted by gestational age per Castillo 2008 reference ranges | Performance | C | Class C: Desat detection failure could result in undetected hypoxemia
DI-003 | UN-003 | Generate JSON-structured handoff reports including patient ID, tier classification, desaturation events, trend summary, and recommended actions | Functional | B | Class B: Report generation supports clinical workflow; failure delays but does not prevent intervention (alarms still active)
DI-004 | UN-004 | Correlate accelerometer signal with SpO2 waveform; flag segments where motion correlation coefficient > 0.6 as artifact-contaminated | Performance | B | Class B: Artifact flagging is supportive; false artifact flag on real data is caught by expert review queue
DI-005 | UN-005 | Calculate classification confidence score (0.0-1.0) per trace; route to expert queue when confidence < 0.7; include top 3 differential classifications | Safety | C | Class C: Expert queue routing failure could suppress ambiguous cases requiring clinical judgment

## Design_Outputs
ID | Linked DI | Output Description | Verification Method
DO-001 | DI-001 | Tier classification module with GA-adjusted threshold lookup table and 4-class output enum | Test
DO-002 | DI-002 | Desaturation detection algorithm with configurable baseline window and duration filter | Test
DO-003 | DI-003 | Handoff report generator producing SBAR-format JSON with all required clinical fields | Inspection
DO-004 | DI-004 | Motion artifact detector using sliding-window cross-correlation with accelerometer channel | Test
DO-005 | DI-005 | Confidence scoring module with softmax-based certainty metric and expert routing logic | Test

## Verification
ID | Linked DO | Test Protocol | Pass Criteria | Status | Test Data Source
VER-001 | DO-001 | TP-001: Run classifier on 500-sample test set with known tier labels | Overall accuracy >= 85%; per-tier F1 >= 0.80 | Pending | Synthetic dataset, N=300 traces, algorithm-generated per GA category
VER-002 | DO-002 | TP-002: Inject synthetic desaturation events at known timestamps into 100 traces | Sensitivity >= 90%; false negative rate < 5% | Pending | Internal NICU reference dataset, N=500 annotated events (Castillo 2008 thresholds)
VER-003 | DO-003 | TP-003: Validate 50 generated reports against SBAR field schema | 100% schema compliance; all required fields present | Pending | Synthetic handoff reports, N=50, template-validated
VER-004 | DO-004 | TP-004: Run artifact detector on 200 segments with labeled motion/no-motion ground truth | Precision >= 80%; recall >= 75% | Pending | Synthetic motion artifact dataset, N=200 traces with accelerometer correlation
VER-005 | DO-005 | TP-005: Verify routing logic on 300 cases spanning confidence range 0.0-1.0 | 100% of cases with confidence < 0.7 routed; zero missed critical (emergency tier) | Pending | Synthetic ambiguous cases, N=100, confidence score distribution validated

## Validation
ID | Linked UN | Validation Method | Acceptance Criteria | Status
VAL-001 | UN-001 | Clinical simulation with 3 neonatologists reviewing 100 classified traces | Clinician agreement with system classification >= 85% | Pending
VAL-002 | UN-002 | Retrospective analysis on NICU dataset with chart-confirmed desaturation events | Sensitivity >= 90% vs. chart-confirmed events | Pending
VAL-003 | UN-003 | Usability study: 10 NICU nurses use handoff reports in simulated shift change | >= 95% task completion; SUS score >= 70 | Pending
VAL-004 | UN-004 | Bench test with motion simulator generating known artifact patterns | Artifact detection aligns with expert annotation in >= 80% of segments | Pending
VAL-005 | UN-005 | Prospective pilot: expert queue reviewed by senior clinician over 2-week period | Zero false negatives for emergency-tier cases; expert queue volume < 20% of total | Pending

## Traceability_Matrix
UN ID | User Need | DI ID | Design Input | DO ID | Design Output | VER ID | VER Status | VAL ID | VAL Status
UN-001 |  | DI-001 |  | DO-001 |  | VER-001 |  | VAL-001
UN-002 |  | DI-002 |  | DO-002 |  | VER-002 |  | VAL-002
UN-003 |  | DI-003 |  | DO-003 |  | VER-003 |  | VAL-003
UN-004 |  | DI-004 |  | DO-004 |  | VER-004 |  | VAL-004
UN-005 |  | DI-005 |  | DO-005 |  | VER-005 |  | VAL-005
Summary Statistics
Verification | Pass: |  | Fail: |  | Pending:
Validation | Pass: |  | Fail: |  | Pending: