## Document_Control
Document Control
Document Title | Risk Analysis — SpO2 AI Eval Pipeline
Document Version | 1.0
Device/Software Version | SpO2 AI Eval Pipeline
Date Generated | 2026-04-28
QMS Document Number
SOP Reference
Related Design Controls File
Approval Signatures
Role | Name | Date | Signature
Prepared By
Reviewed By (QA)
Approved By (Management)

## Hazard_Identification
Hazard ID | Hazard | Hazardous Situation | Harm | Sequence of Events
HAZ-001 | Misclassify urgent desaturation as normal | Algorithm classifies a true urgent desaturation event as normal SpO2, and no alarm is triggered | Delayed clinical intervention, potential brain injury | Sensor captures low SpO2 -> AI model misclassifies severity -> no alarm fired -> clinician unaware -> delayed O2 therapy -> hypoxic brain injury
HAZ-002 | False positive artifact flag on real desaturation | Real desaturation event is flagged as motion artifact and discarded from analysis | Real desat discarded, delayed clinical response | True desat occurs -> artifact filter misidentifies as noise -> event dropped -> clinician not notified -> delayed response
HAZ-003 | Sensor disconnect undetected | Pulse oximetry sensor becomes disconnected and system does not alert nursing staff | No monitoring during disconnect period | Sensor dislodged -> signal loss -> system does not detect disconnect -> monitoring gap -> event missed
HAZ-004 | Wrong GA threshold applied | Gestational age category is incorrectly assigned, causing wrong SpO2 thresholds to be used for classification | Inappropriate urgency classification | GA entered incorrectly or defaulted -> wrong threshold table selected -> classification uses wrong limits -> over/under-triage
HAZ-005 | Handoff report omits critical context | AI-generated shift handoff report is missing key clinical context needed for care continuity | Clinician makes suboptimal decision due to missing context | Handoff report generated -> critical event omitted -> receiving clinician unaware of history -> suboptimal care decision

## Risk_Estimation
Hazard ID | Severity (S1-S5) | Probability (P1-P5) | Risk Level
HAZ-001 | 4 | 3
HAZ-002 | 3 | 4
HAZ-003 | 3 | 3
HAZ-004 | 3 | 2
HAZ-005 | 2 | 3

## Risk_Evaluation
Hazard ID | Risk Level | Acceptability | Rationale
HAZ-001 |  |  | Residual risk is ALARP: dual-threshold with GA adjustment reduces probability to Incredible; further reduction requires fundamentally different sensor technology which is impracticable
HAZ-002 |  |  | Residual risk is ALARP: confidence scoring routes uncertain cases to expert review; residual probability is Improbable
HAZ-003 |  |  | Residual risk is Acceptable: continuous signal check reduces probability to Incredible
HAZ-004 |  |  | Residual risk is Acceptable: validation at admission with boundary checks reduces probability to Incredible
HAZ-005 |  |  | Residual risk is Acceptable: mandatory fields ensure critical data inclusion; probability reduced to Improbable

## Risk_Controls
Control ID | Hazard ID | Control Measure | Type | Verification Method
RC-001 | HAZ-001 | Dual-threshold validation with GA-adjusted limits | Inherent safety by design | Unit tests with edge-case SpO2 waveforms; clinical validation against expert labels (n>=200)
RC-002 | HAZ-002 | Confidence scoring with expert queue fallback | Protective measure | ROC analysis on artifact vs. real desat test set; expert review queue audit (monthly)
RC-003 | HAZ-003 | Signal continuity check every 5 seconds | Protective measure | Simulated disconnect test (n=50); verify alarm triggers within 10s of signal loss
RC-004 | HAZ-004 | GA category validation at admission | Inherent safety by design | Input validation unit tests; GA boundary-value testing; admission workflow audit (quarterly)
RC-005 | HAZ-005 | Structured template with mandatory fields | Information for safety | Template completeness validation; clinician usability testing (n>=10); monthly content audit

## Residual_Risk
Control ID | Hazard ID | Post-Control Severity (S1-S5) | Post-Control Probability (P1-P5) | Residual Risk Level | Rationale
RC-001 | HAZ-001 | 4 | 1 |  | Residual risk is ALARP: dual-threshold with GA adjustment reduces probability to Incredible; further reduction requires fundamentally different sensor technology which is impracticable
RC-002 | HAZ-002 | 3 | 2 |  | Residual risk is ALARP: confidence scoring routes uncertain cases to expert review; residual probability is Improbable
RC-003 | HAZ-003 | 3 | 1 |  | Residual risk is Acceptable: continuous signal check reduces probability to Incredible
RC-004 | HAZ-004 | 3 | 1 |  | Residual risk is Acceptable: validation at admission with boundary checks reduces probability to Incredible
RC-005 | HAZ-005 | 2 | 2 |  | Residual risk is Acceptable: mandatory fields ensure critical data inclusion; probability reduced to Improbable

## Overall_Residual_Risk
Category | Description | Status | Notes
Individual Residual Risks | All individual hazard residual risks are documented in the Residual_Risk sheet
Risk Interaction Analysis | Assessment of whether residual risks interact or compound | Reviewed | HAZ-001 (misclassification) and HAZ-002 (false artifact) could compound: if artifact flag suppresses real desat AND classification misses it, double failure path exists. Mitigated by expert queue (RC-005) which catches both.
Risk Clusters Identified | Groups of related residual risks | Documented | Cluster 1: Information hazards (HAZ-001, HAZ-002, HAZ-005) — all relate to incorrect clinical information. Cluster 2: Detection hazards (HAZ-003, HAZ-004) — relate to system configuration/connectivity.
Benefit-Risk Conclusion | Overall assessment of acceptability considering intended use benefits | Acceptable | Benefits of automated SpO2 triage (reduced alarm fatigue, faster clinical response, GA-adjusted thresholds) outweigh residual risks. All unacceptable risks reduced to ALARP or Acceptable. No unacceptable residual risks remain.
Post-Market Monitoring Plan | How field data feeds back into risk analysis | Defined | Complaint handling per QMS procedure. Quarterly review of classification accuracy vs. clinical outcomes. CAPA process triggers re-analysis if new failure modes identified.
Overall Conclusion

## FMEA
Hazard ID | Failure Mode | Effect | Cause | Severity (S) | Occurrence (O) | Detection (D) | RPN
HAZ-001 | Classification algorithm outputs 'normal' for SpO2 < 85% | No alarm triggered for urgent desaturation | Model undertrained on preterm-specific waveforms | 8 | 4 | 3
HAZ-002 | Artifact filter rejects waveform segments with SpO2 < 88% | Desaturation event not recorded or alarmed | Artifact filter threshold too aggressive for low-signal conditions | 6 | 5 | 4
HAZ-003 | System does not generate disconnect alarm within 30s | Monitoring gap during sensor disconnect | Signal loss detection logic has timeout > clinical threshold | 6 | 3 | 2
HAZ-004 | GA lookup returns incorrect category for edge-case gestational ages | SpO2 thresholds do not match patient's clinical profile | Off-by-one error in GA week boundary logic | 6 | 3 | 2
HAZ-005 | Handoff template renders with empty mandatory fields | Clinician receives incomplete handoff information | Template logic does not enforce required field population | 4 | 4 | 3